## Slide 1
CAPSTONE PROJECT DEFENSE
Edge-AI Automated Peripheral Blood Smear Analysis
A YOLOv8n Deployment on Raspberry Pi 4 for Resource-Constrained Hematology Screening
Shreyash Dhoot | Abhishek Karad | Pranav Lute | Prince Gupta
Project Guide: Dr. Anagha Deshpande, Assistant Professor
Department of Electrical and Electronics Engineering (DOEEE)
MIT World Peace University (MIT-WPU), Pune, Maharashtra, India

## Slide 2
Clinical Context: The Diagnostic Gap
Edge-AI Blood Smear Analysis | MIT-WPU DOEEE
2
Hematology testing is the gateway to diagnosing anemia, leukemia, sepsis and thrombocytopenia, yet access is limited where it is needed most.
$30k - $150k
Automated CBC analyzers, plus proprietary reagents, stable power and technicians
5 - 15 min
Manual 100-cell differential per slide; subjective, with human-vs-human ICC of only 0.45 - 0.60
24 - 72 h
Typical wait as samples travel to district reference laboratories
CORE OBJECTIVE
A battery-powered, point-of-care, $275 automated cytology system that delivers leukocyte and platelet triage in under 35 seconds per patient.
Notes: Hematology testing is the gateway to diagnosing anemia, leukemia, sepsis and thrombocytopenia. Automated analyzers are absent in primary health centers, rural clinics and disaster zones. Manual smear reading is the gold standard fallback but is slow and subjective.

## Slide 3
Hardware Architecture and Optical Setup
Edge-AI Blood Smear Analysis | MIT-WPU DOEEE
3
1
Computing core
Raspberry Pi 4B: BCM2711, quad-core Cortex-A72 @ 1.5 GHz, 4 GB LPDDR4
2
Optical sensor
12.3 MP Sony IMX477 back-illuminated CMOS, 1.55 µm pixel pitch
3
Mechanical interface
Custom 3D-printed C-mount adapter coupled to a 23.2 mm trinocular microscope tube
4
Microscope
Standard compound laboratory scope, 100× oil immersion, Giemsa-stained smear
TOTAL CAPEX
$275
Single-board computer $75
12.3 MP sensor $50
Optical frame $150
99% cheaper than commercial automated analyzers

## Slide 4
Embedded Deep Learning Pipeline
Edge-AI Blood Smear Analysis | MIT-WPU DOEEE
4
1
PyTorch training
best.pt
6.2 MB
2
ONNX export
Fixed 640 × 640 input
11.7 MB
3
INT8 quantization
Dynamic quantization
3.2 MB
4
ONNX Runtime
CPU provider, all 4 ARM cores
Deployed
YOLOv8n (nano): anchor-free single-stage detector with a CSPDarknet backbone, Path Aggregation Network neck, decoupled detection heads and 3.2M parameters.
245 ms
latency per frame
4.1 FPS
live viewfinder rate
3.4 W
sustained power draw
Notes: Why YOLOv8n and not UNet or Mask R-CNN: edge latency was the primary design constraint. A field has over 100 cells; Mask R-CNN or UNet would need over 2,500 ms per frame on the Pi's ARM CPU, precluding a live viewfinder. YOLOv8n gives anchor-free box regression in 245 ms with 3.2M parameters, while keeping 0.98 F1 on leukocytes.

## Slide 5
In-Domain Performance: BCCD Benchmark
Edge-AI Blood Smear Analysis | MIT-WPU DOEEE
5
Chart: Per-class detection metrics (IoU 0.5)
| WBC | Platelets | RBC
Precision | 0.954 | 0.783 | 0.496
Recall | 0.997 | 0.917 | 0.958
F1-score | 0.975 | 0.844 | 0.654
AP@0.5: WBC 0.985 | Platelets 0.834 | RBC 0.749
0.856
mAP@0.5
(mAP@0.5:0.95 = 0.617)
BCCD dataset
364 images at 100× magnification; 4,155 RBCs, 372 WBCs, 361 platelets
Near-flawless WBC identification and high platelet precision on standardized smears.

## Slide 6
Clinical Agreement (CLSI EP09-A3)
Edge-AI Blood Smear Analysis | MIT-WPU DOEEE
6
White blood cells
+0.047
mean bias, cells per field
Bland-Altman limits of agreement [-1.56, +1.65] cells
Spearman ρ = 0.771 (p < 0.001)
Platelets
0.862
ICC(2,1) agreement
Deming regression slope 1.012, near identity
Bias -0.456; limits [-4.20, +3.28]
Red blood cells
0.112
ICC(2,1) agreement
Recall is high (0.96), but count agreement is low
Overlap at ~700:1 density gives 4,042 partial-edge false positives
Evaluated against CLSI EP09-A3 method-comparison standards, not just detection mAP.
Notes: Why the RBC ICC is only 0.112 despite 0.96 recall: red cells outnumber white cells about 700 to 1. Dozens intersect the image border or overlap in each field, and bounding-box detectors produce 4,042 partial-edge false positives. Leukocytes have dense chromatin nuclei that remove edge ambiguity. Future RBC enumeration will use area-density regression instead of discrete boxes.

## Slide 7
Out-of-Distribution Test: Clinical 72
Edge-AI Blood Smear Analysis | MIT-WPU DOEEE
7
Chart: F1-score under domain shift
| WBC F1 | Platelet F1
BCCD (in-domain) | 0.98 | 0.84
Clinical 72 (zero-shot) | 0.82 | 0.17
72 real clinical smears from an independent hospital laboratory, evaluated zero-shot.
Stain duration, buffer pH and lighting temperature shift Romanowsky dye appearance.
Leukocytes stayed robust: F1 0.82, precision 0.958, AP@0.5 0.691.
Platelets dropped sharply: recall 0.123, F1 0.173, AP@0.5 0.057.
Platelets lack dense nuclei, so focal drift and color changes easily mask them.
Notes: Zero-shot evaluation on 72 real smears from an independent hospital laboratory. Leukocytes stayed robust; platelets collapsed. Different stain duration, buffer pH and lighting alter Romanowsky dye appearance.

## Slide 8
Domain Adaptation Pipeline
Edge-AI Blood Smear Analysis | MIT-WPU DOEEE
8
1
Invariance-augmented training
180° rotation, vertical flips (p = 0.5), MixUp regularizer 0.15, dropout 0.1
2
Reinhard stain normalization
Aligns image color mean and standard deviation in CIELAB space to a standardized BCCD reference smear
3
CLAHE
Contrast-limited adaptive histogram equalization on the L channel (clip limit 2.0, 8 × 8 grid) to sharpen cell boundaries
4
Lineage-specific thresholding
Decoupled decision boundaries: τ(PLT) = 0.15 and τ(RBC/WBC) = 0.25

## Slide 9
Generalization Ablation Study
Edge-AI Blood Smear Analysis | MIT-WPU DOEEE
9
Configuration | Split | Preprocessing | Post-processing | mAP@0.5 | RBC F1 | WBC F1 | PLT F1
(a) Baseline YOLOv8n | BCCD test, n=364 | None | Default τ=0.25 | 0.856 | 0.65 | 0.98 | 0.84
(b) Baseline zero-shot | Clinical 72, n=72 | None | Default τ=0.25 | 0.410 | 0.63 | 0.82 | 0.17
(c) Retrained v2 | Clinical 72, n=72 | None | Default τ=0.25 | 0.411 | 0.67 | 0.82 | 0.06
(d) Retrain + Stain Norm | Clinical 72, n=72 | Reinhard + CLAHE | Default τ=0.25 | 0.398 | 0.66 | 0.82 | 0.03
(e) Retrain + Thresholding | Clinical 72, n=72 | Reinhard + CLAHE | τ(PLT)=0.15 | 0.414 | 0.66 | 0.82 | 0.21
(f) Retrained in-domain | BCCD test, n=364 | None | Default τ=0.25 | 0.829 | 0.64 | 0.98 | 0.84
Highlighted row (e): the only configuration that recovers platelets under stain normalization.
Notes: Walk through rows (a) to (f). Row (d) shows stain normalization alone suppressing platelets; row (e) is the only configuration that recovers them, via per-class thresholding.

## Slide 10
The Platelet Thresholding Rescue
Edge-AI Blood Smear Analysis | MIT-WPU DOEEE
10
The paradox
Reinhard normalization alone dropped platelet F1 to 0.03. It weakened the faint pale-purple hue of thrombocytes, pushing detector confidence just below the default 0.25 threshold.
The fix
Calibrating the platelet decision boundary to τ(PLT) = 0.15 (row e) rescued recall and lifted F1 back up, a 6.4× recovery.
0.018 → 0.175
platelet recall
0.03 → 0.21
platelet F1
6.4×
recovery
Core rule: color normalization shifts head activations, so couple every preprocessing transform with threshold recalibration.
Notes: Platelets are 2 to 3 µm, nucleus-free and faintly pale-purple. Reinhard normalization aligns global channel statistics in CIELAB and diluted their chromatic contrast, pushing raw confidence scores just under 0.25. Lowering tau(PLT) to 0.15 recovered recall from 0.018 to 0.175. Normalization must always be paired with threshold recalibration.

## Slide 11
Edge Systems, Power and Economics
Edge-AI Blood Smear Analysis | MIT-WPU DOEEE
11
Commercial 3-part analyzer
$30,000
plus $2,500 per year in reagents
Our Edge-AI microscope station
$275
one-time cost, zero proprietary reagents
3.67×
model compression
11.7 MB → 3.2 MB
4.1 FPS
245 ms per frame
within operator scan speed
3.4 W
sustained draw
>5 h on a 20 Wh power bank
< 58 °C
core temperature
aluminum heatsink, no fan
INT8 quantization keeps WBC F1 unchanged at 0.98 while cutting model size by 72.7%.

## Slide 12
Point-of-Care Clinical Workflow
Edge-AI Blood Smear Analysis | MIT-WPU DOEEE
12
1
Prepare the slide
Fingerprick drop, wedge smear, 3-minute field stain or Giemsa dip
2
Mount and focus
Place the slide on the stage and focus under 100× oil immersion
3
Live view
Camera streams a 4.1 FPS viewfinder to a 7-inch touchscreen
4
Automated differential count
Scan 10 - 20 fields; counts accumulate and abnormal clusters are flagged
25 - 35 s
digital triage report per patient
TRIAGE BOUNDING
Triages acute infections and leukopenia on the spot
Flags suspicious smears for secondary pathologist review

## Slide 13
Limitations and Threats to Validity
Edge-AI Blood Smear Analysis | MIT-WPU DOEEE
13
!
Platelet OOD sensitivity
Cross-center platelet recall requires on-site staining calibration.
!
Erythrocyte overlap
Bounding boxes overcount overlapping RBCs; contour watershed segmentation is needed for precise hematocrit estimation.
!
3-class, not 5-part differential
The system detects general WBCs; splitting into neutrophils, lymphocytes, monocytes, eosinophils and basophils needs higher optical resolution.
!
Validation cohort size
The clinical split (n = 72) is single-center archival slides; multi-center trials are planned.

## Slide 14
Conclusion and Future Roadmap
CONTRIBUTIONS DELIVERED
A functional $275 Edge-AI cytological screening station
INT8-quantized YOLOv8n at 4.1 FPS on a 3.4 W Raspberry Pi 4
Edge hematology evaluated against CLSI EP09-A3 clinical standards
Ablation linking Reinhard stain normalization to lineage thresholds
NEXT STEPS
Automated 5-part leukocyte differential model
Motorized 3D-printed stage with stepper motors for automated serpentine slide scanning
Direct area-density regression for precise red blood cell enumeration
14
Notes: Can the system diagnose leukemia today? No, and we explicitly avoid that claim. It is a triage screener for primary care: it flags severe infection, leukopenia or leukocytosis in under 35 seconds and directs the patient to a reference hospital for confirmatory flow cytometry.

## Slide 15
Questions?
Open for questions from Dr. Anagha Deshpande and the evaluation committee.
Shreyash Dhoot | Abhishek Karad | Pranav Lute | Prince Gupta
DOEEE, MIT World Peace University, Pune